Register branches to aortic arch even if branch detection fails

MGH Sheep Protocol
Reach out to Brian before the study and make sure the dash monitor is free!!

Materials checklist:
Book relevant materials on google calendars
Computational cart:
- IVUS machine
- EM station
- EM emitter
- Mitra clip
- Anthony lab EM tracker stand
- PIM module
- USB splitter
- pullback device
- pullback device power chord
- pullback device USB cable
- dash monitor cable
- computer
- computer monitor facing out
- computer charger
- frame grabber
- optional: frame grabber for fluoro
- mouse
- keyboard
- supply boxes
- dash monitor
- camera
- camera stand
- iPhone stand

supplies:
- printed protocol
- barbell weights
- IVUS + EM catheters
- steerable EM catheter
- steerable catheter
- computer monitor
- guiding catheters
- dryseal
- dilators
- hemostatic valve
- guidewires
- contrast agent
- saline
- stents
- puffing catheters
- surgical drapes
- ECG electrodes
- syringes for contrast agent injection

Room Setup:

Electronics setup:
- setup power strip
- power on IVUS
- set up EM emitter stand, power on EM station
- plug in computer charger and usb-splitter

Pullback device setup:
- Ask staff if MGH dash monitor can be used for monitoring ECG
- setup pullback device on tripod
- connect ECG electrodes to dash monitor
- connect ECG electrodes to animal
- connect dash monitor to pullback device

Personnel:
- Tom – running ultrasound platform
- Brian – performing procedure
- Diego – repositioning fluoro machine, hand off of tools

EM tracker stand setup:
- Position EM tracker at level of pig’s abdomen (where we want tracking to be most accurate)
- Place a barbell on EM tracker stand
- Give warning to team and animal facility staff not to move EM tracker stand

Camera setup:
- Ensure iPhone is set to record at 60fps
- Ensure fluoro machine is in record mode / connect a frame grabber
- Station camera with view of operator, C-arm, ultrasound 3D display, ECG signal, pullback device, and possibly EM tracker stand
- station other camera / iPhone to record fluoro screen

Computer setup:
- ensure correct calibration file loaded for catheter
- ensure correct ML model loaded
- presave data directories on computer
- activate frame grabber, ensure aortascope working
- (optional) connect frame grabber to fluoro machine

Protocol:

Gaining access:
- Use vessel puncture kit to gain femoral access
- Feed 0.014” wire in (compatible with steerable system)
- Exchange puncture kit for 16Fr dilator inside a 16Fr sheath
  - if that won’t go into the vessel, try 12 Fr (though won’t be able to use steerable system or 0.035” from that point on)
- Remove 16Fr dilator
- Feed in IVUS 0.018” over 0.014” wire
- Check initial imaging quality and tune gains as necessary
  - If vessel diameter too big at first glance, switch to 0.035” IVUS inside 16Fr only for rest of the study – will not be able to use steerable sheath but may still be ok in a sheep
- Push the catheter up to the base of the aortic arch, and do a quick pullback data collection of the abdominal aorta to ensure at least one dataset gets collected – see if the registration quality is good enough

Navigating aortic arch – approach 1 (easy approach) – use precurved guiding catheter
- Feed a guiding catheter over the wire and navigate it up to the base of the aortic arch
- Using fluoroscopy as guidance, rotate the guiding catheter to point in the direction of the aortic arch and pass an 0.018” or 0.014” wire through until it reaches the base of the aortic root
  - inject a puff of contrast if aortic root direction is unclear, take a picture for paper
- Remove the guiding catheter
- Test to see if the IVUS catheter can be feed around the aortic arch with minimal force without entering the head branches

Navigating aortic arch – approach 2 (easy approach) – use Mitra Clip
- Push dryseal sheath up to base of aortic arch
- Deploy mitra clip from sheath and steer around arch, or feed wire out of mitra clip if you have to after point in right direction (take video of mitra clip navigation)
- Feed IVUS through and retract mitra clip
- Retract everything together during the pullback
Navigating aortic arch – approach 2 (if approach 1 doesn’t work) – use steerable catheter
- Remove IVUS 0.018”
- Feed 12 Fr sheath over 0.014” inside 16Fr sheath
- Navigate 0.014” wire around aortic arch
- Feed steerable catheter over 0.014” inside 12Fr
- Navigate steerable system around aortic arch by having it follow the wire
- Once the steerable system is at the aortic root, bend the 12Fr sheath around the arch by pushing it over the steerable system
- Once sheath is at aortic root, remove the steerable inner component
- Feed IVUS up through the sheath

Pullback (repeat for each dataset collection):
- Ensure gating is on and other yaml parameters look correct, check whether fluoro arm creates EM interference (when both on and off)
- Remind staff in room not to touch EM tracker stand from this point onwards
- Create an empty dataset folder for pullback data and load preop data into it
- Plug IVUS catheter in and ensure adequate ultrasound imaging quality before recording and that catheter is outside the sheath
- Start recording with (1) camera in room, (2) screen recording, (3) click ‘start recording’ in aortascope, and (4) click start pullback device in that order
- (for one experiment) Take a picture and video of OR setup that includes – EM tracker stand, C-arm, ECG signal, pullback device, operator, ultrasound 3D display
- Pull to the end of the distal abdominal aorta smoothly and continuously, (1) click ‘save data’, then (2) ‘stop pullback device’.
  - If you used the steerable system, pull the whole system back together (catheter + sheath) with no relative displacement between the components
- Flush IVUS catheter with saline
Registration contingency plans
- Have deformation registration directory also open for troubleshooting
- Redo pullback
- Record transform data and use IVUS mapping only (superimpose CT scan later)
- Branch click function
- Repeat with abdominal registration
- Repeat with refine registration if arch map was good
- Try just running registration again
- Load a different ML model – you can click replay dataset
- Try running viterbi, nearest neighbour, and fenfit
- Turn gating on

Side branch navigation:
- Turn ‘record_poses’ parameter on
- Remove IVUS catheter and guidewire
- Position the catheter so that it includes a view of all abdominal branches and the distal abdominal aorta, use contrast as needed
- Remind staff in room not to touch EM tracker stand
- Plug in steerable catheter, launch AortaScope, switch endoscopic view
- Navigate to bottom of distal abdominal aorta, just above the femorals
- Start recording with (1) camera in room, (2) screen recording
- Start the clock – interventionalist attempts to cannulate Celiac, SMA, LRA, and RRA as quickly as possible from top to bottom
  - Ultrasound view is continuously manipulated to be at 90 degrees to the vessel for clear side view
  - perform a shot of fluoroscopy to verify successful cannulation (quick puff of contrast through the steerable catheter on ONE of the cannulations, taking a video while doing so)
  - Record time each time a vessel is cannulated, pausing timer when cannulation has been verified to reset (use ‘lap’ function on iPhone and take screenshot at end of study)
  - Record the fluoro use time for each vessel
  - Record time in datasheet when 90 degree view of the next vessel is achieved
- (for one experiment) Take a picture and video of OR setup that includes – EM tracker stand, C-arm, ECG signal, pullback device, operator, ultrasound 3D display
- Flush steerable catheter with saline after each experiment
- (whenever CAC approved / ex vivo) on the last cannulated vessel, feed balloon stent (Medtronic Integrity) over cannulated wire and deploy under fluoro, recording fluoro screen as you navigate stent and deploy – inject a puff of contrast after if possible

Fluoro comparison
- Leave fluoro machine in same spot
- Using same catheter, attempt to cannulate all vessels beginning at the distal abdominal aorta
- Rotate C-arm as necessary, including rotation in required cannulation time
- Record the fluoro use time for each vessel
- (for one experiment) Take video of surgeon rotating C-arm to gain 90 degree view of vessel while trying to cannulate and injecting contrast at same time