IVUS-guided Aortic Branch Cannulation – in vitro protocol

Timetable:
Option 1 - all studies at MIT (5 hour long study):
- Each phantom should take 10 mins per user (5 mins max to do cannulation with ultrasound + 5 mins max to do cannulation with fluoroscopy)
- for 4 users - that’s ~40 mins per phantom
- Each phantom should take 15 mins to map
- split groups roughly in half and rotate whether X-ray or ultrasound is done first each time (Group 1 (G1) – Brian, Ben Group 2 (G2) – Brett, Chase)
- Things to point out when participants arrive - point to lead, start at bifurcation, how to hit branch and reference point on the wire, verifying cannulation with fluoro, 7 phantoms, things we will measure, group orders, timetable, VPA forms, warning to not touch EM sensor / bed, show steerable catheter
- If short on time → do IVUS only, schedule with Brian / Ben only another time

Patient K6 mapped beforehand – make sure good view

12:00 Patient K6 study starts – G1 (Ultrasound then Xray); G2 (Xray then Ultrasound)

12:40 Patient K1 is mapped (pump on)
12:55 Patient K1 study starts - G1 (Xray then Ultrasound); G2 (Ultrasound then Xray)

1:35 Patient K12 is mapped (pump on)
1:50 Patient K12 study starts - G1 (Ultrasound then Xray); G2 (Xray then Ultrasound)

2:30 Patient K5 is mapped (pump on)
2:45 Patient K5 study starts – G1 (Xray then Ultrasound); G2 (Ultrasound then Xray)

3:20 (10min) Patient K3 is mapped (pump on)
3:30 (30 min) Patient K3 study starts - G1 (Ultrasound then Xray); G2 (Xray then Ultrasound)

4:00 (10min, PVA) Patient K4 is mapped
4:10 (30 min, PVA) Patient K4 study starts - G1 (Xray then Ultrasound); G2 (Ultrasound then Xray)

4:40 (10 min, PVA) Patient K8 is mapped
4:50 (30 min, PVA) Patient K8 study - G1 (Ultrasound then Xray); G2 (Xray then Ultrasound)

Room Setup:

- release air from top with syringe [ ]
- test pullback device [ ]
Give warning to user to not move tank or EM emitter [ ]
- Ensure iPhone is setup to record at 60fps majority view [ ]
- Ensure Nikon camera is set to 25 fps, long lens, place very up close on high stand [ ]
- ensure correct calibration file loaded for catheter [ ]
- ensure correct ML model loaded [ ]
- presave data directories on computer for all 7 phantoms and load preop data for each! [ ]
- Have deformation registration directory also open for troubleshooting [ ]
- Ensure gating is on and other yaml parameters look correct [ ]
- Turn pump frequency down to ~70bpm [ ]
- Use syringe to remove air from pulsatile pump [ ]
- ensure tubings submerged where possible [ ]

Mapping of phantom (for each phantom):
- Place phantom in tank, removing bubbles
- Turn pump on and tune flow speed by changing constriction at the suction point
- (Inkbit) Feed 16 Fr sheath up to level of aortic arch, feed guidewire from sheath to aortic root
- Feed IVUS 0.035 catheter to aortic root
- (PVA) Clip IVUS catheter into pullback device
- Plug IVUS catheter in and ensure adequate ultrasound imaging quality before recording and that catheter is outside the sheath – tune gains as necessary
- Create an empty dataset folder for pullback data and load preop data into it

Pullback data collection each time:
- Ensure data directory with preop data available
- Start recording with (1) camera in room, (2) screen recording, (3) click ‘start recording’ in aortascope, and (4) click start pullback device in that order
- Pull to the end of the distal abdominal aorta smoothly and continuously, (1) click ‘save data’, then (2) ‘stop pullback device’.
- (inkbit) pull the whole system back together (catheter + sheath) with no relative displacement between the components
- (For 4D PVA motion) gating, replay bin 3 and 8, register to bin 3, initialize cardiac motion, load

Registration contingency plans (20 mins)

1. Run registration again – or try viterbi
2. Replay data – conf class 2 threshold, tsdf voxel size, ML model
3. Registration tuning – dbscan eps, downsampling CT mesh, allowable stretch factor, angle filter ransac, renal order swap injective mappings, bin / full selected for 4D reg
4. IVUS mapping only and record transform data (superimpose CT scan later)

(worst case) Redo pullback

Side branch navigation:

Setup:
- Position the fluoroscopy view so that it includes a view of all abdominal branches and the distal abdominal aorta - Don’t move the C-arm from here! Otherwise magnetic field will change
- Cover phantom with black cloth, plug in steerable catheter

Ultrasound-guided aortic branch cannulation:
0. Save previous surgeon’s transform data (if possible)
0. Turn on pulsatile pump – for video even if pullback device not used
1.Position catheter at bottom of distal abdominal aorta, just above the femorals (use both fluoro and aortascope to verify)
2. Turn on camera recording with (1) camera in room, (2) screen recording
3. Start the clock and have interventionalist attempt to cannulate:
- During the study, the ultrasound view is continuously manipulated to be at 90 degrees to the side branch vessel for clear side view
- When the surgeon thinks they have cannulated, pass the wire in, and perform a shot of fluoroscopy to verify successful cannulation – test puffing through steerable (video of screen)
- Record time it takes to cannulate each vessel (use ‘lap’ function on iPhone, pause timer once vessel is cannulated to allow reset)
- Record the total time that fluoro is used for during the study
4. Every time surgeon fires fluoro = 1 cannulation attempt

Comparison to fluoroscopy:
- Record video of fluoro screen and surgeon
- Any time surgeon pushes wire – cannulation attempt – draw a strike on sheet

Materials checklist:
Remember to book relevant materials on lab google calendars

Computational cart:
- IVUS machine
- EM station
- EM emitter
- EM tracker stand (Anthony lab)
- PIM module
- USB splitter
- marker for drawing on fluoro screen
- basin to prevent water leakage
- pump-aorta large inlet connector
- black cloth
- second cart for monitor (for now)
- pullback device
- pullback device power chord
- pullback device USB cable
- pulsatile pump
- computer
- computer monitor facing out
- computer charger
- frame grabber
- optional: frame grabber for fluoro
- optional: additional stand for computer monitor
- mouse
- keyboard
- supply boxes
- dash monitor
- camera
- camera stand
- iPhone stand
- data recording sheets
- pens
- screwdriver
- scissors
- lighter
- beaker
- extra basin for dumping water
- wooden/plastic board as metal table alternative
- barbell weights
- guidewires consolidated in box sizes
- loop closure connectors
- PVA phantom boxes
- inkbit phantoms
- M6 wrench
- towel
- dryseal
- velcro (male and female)

supplies:
- printed protocol
- barbell weights
- IVUS + EM catheters
- steerable EM catheter
- steerable catheter
- computer monitor
- guiding catheters
- guidewires
- contrast agent
- puffing catheters
- surgical drapes
- syringes for contrast agent injection